Food and Pharmacy industries RTD temperature probes（TM1601)

Applications:

Food and Pharmacy industries etc.

Technical data

Protection tube：stainless steel 304/316

Terminal head: stainless steel 316

Sensor element:  Pt100 meets IEC 751

Tolerance :     Class A: △t=±(0.15 + 0.002 |t| )℃(standard)

1/3Class B: △t=±(0.10 + 0.0017 |t| )℃

Protection Class： IP65

Transmitter:Output 4-20mA, 2-wire Power supply: 12-30VDC

Load resistance 0-600Ω/24V

Accuracy:±0.5%FS，±0.2%FS(optional）

接线图-Wiring diagram 铂电阻RTD

单支2线 1x2-wire单支3线 1x3-wire单支4线 1x4-wire

双支4线 2x4-wire双支6线 2x6-wire红色标记 red mark

安装螺钉 mounting screw

Order Details:

| TM1601 | — | (1) | — | (2) | — | （3） | — | (4) | — | (5) | — | (6) | / | (7) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |

- (1)protection tube diameter D

6      6mm

8      8mm

9      9mm

11      11mm

- Insertion length EL

- 50     50mm

- 100   100mm

- 150   150mm

- 200    200mm

- 250    250mm

Note: other lengths as per customers’ requirements(EL≤1000mm）

- Process connection G

213       G1/2 thread

410       DN10  DIN11851

411       DN25  DIN11852

412       DN32  DIN11853

413       DN10/20 DIN32676

414       DN25/40 DIN32677

415       DN50 DIN32678

416       DN10/15 VARIENT

417       DN25/32 VARIENT

418       DN40/50 VARIENT

510       Spherical welding units

- Details in “Thread code table”

- Sensor element

- 11      1 x Pt100 in 2-wire circuit

51      1 x Pt500 in 2-wire circuit

- 61      1 x Pt1000 in 2-wire circuit

13      1 x Pt100 in 3-wire circuit

14      1 x Pt100 in 4-wire circuit

12      2 x Pt100 in 4-wire circuit

16      2 x Pt100 in 6-wire circuit

62      2 x Pt1000 in 4-wire circuit

(5)Measuring range

14    -50℃ ～+ 260℃

(6) Cable length L（L≥100mm）

100     100mm

200     200mm

2000    2000mm

Note: other lengths as per customers’ requirements

(100mm≤L<1000mm, L1=50mm, L≥1000mm, L1=100mm)

(7)Additional options

00      No

02      shielded connecting cable

06      Output 4-20mA in 2-wire

07      Ferrules as per DIN 46 228 part 4(standard)

08      Multipole connector (providing types)

99      Special options

List the additional options in sequence, separated by commas.

Dimensions

锥台密封螺纹连接 符合CIP要求

cone sealing threaded connections.

Meets the requirements as per CIP.

标准卡箍连接-- standard clamp connection

球型焊座 带夹紧装置 --ball type welding base. With clamping devices.

DIN 11 851 标准

带环形螺母的锥形管接头

DIN 11 851 standard

Tapered pipe fittings with ring nut.

焊座 welding base

密封系统与CPI兼容  sealing system is compatible with  CPI

球型焊套-- ball welding thermowell.